[image: image1.jpg]DELHI POLLUTION CONTROL COMMITTEE
DEPARTMENT OF ENVIRONMENT: GOVT. OF NCT OF DELHI
4™ & 5™ FLOOR: ISBT BUILDING: KASHMERE GATE: DELHI-6

F.No.DPCC/Lab (W) /2010/ )7 J? Date:- / l,/ 9///

E-TENDER NOTICE

Sub:- Quotation forjurchase of 02 (I'wo) nos. of Stack Monitoring Kit for DPCC
Laboratory.

As directed by directed by competent authority, 1 am informed to invite the quotations
on-line through E-Tender website i.e. https://govtprocurement.delhi.gov.in for purchase of 02
(Two) nos of Stack Monitoring Kit for DPCC laboratory only from manufacturer or their
authorized agents. Specification of Stack Monitoring Kit is attached herewith as “Annexure-
A?,

A. DETAILS OF TENDER

a) Last date of downloading of Tender Documents: - 10.10-2011 at 12.00 Noon.
b) Last date of submitting of Tender Documents: - 10-10-2011 at 14.00 Noon.
¢) Opening of Technical Tender :- 11-10-2011 at 15.30 PM

d) Opening of Financial Tender :- 11-10-2011 at 16.00 PM

B. Eligibility Criteria:-

Following Essential Documents should be submitted with the technical bid failing
which the tender shall be rejected.

1. Only manufacturer or their authorized agents are eligible to participate in the
quotations. Valid supporting documents regarding manufacturer or authorized

agents shall be attached.
\9\0 2. Earnest Money of Rs. 5,000/- (Five Thousand Only). in the form of DD/FD

Q/ receipt in favour of “Delhi Pollution Control Committee” payable at Delhi.
/_)((1.7 P ¢ 3. Self Attested copy of VAT / TIN Number.
A 4. Technical details & information brochures of Stack Monitoring Kit (printed
\\)‘ product brochures and all technical leaflets, etc) to be attached compulsorily
with the technical offer / bids.
5. The tenderer shall also upload the scanned copies of above documents on E-
Tender website i.e. https:/govtprocurement.delhi.gov.in failing which the
tender shall be rejected.
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1 The prices must be quoted inclusive of all taxes & statutory levies and charges i.e.
basic price, taxes, packing, forwarding & transportation charges, insurance,
installation charges etc.

2 Sealed Technical quotation must be submitted physically in Delhi Pollution Control
Committee, 4th Floor, ISBT Bldg., Kashmere Gate, Delhi-110006 before or on
10.10.2011 at 3.00 PM and be drqpped in quotation box placed at Enquiry Counter,.
The sealed technical envelop shall be superscribed as “ Technical quotation for Stack
Monitoring Kit” and addressed to Member Secretary, Delhi Pollution Control
Committee, 4" Floor, ISBT Building, Kashmere Gate, Delhi-110006.

3. Only manufacturer or their authorized agents are eligible to participate in the
quotations.
4. Minimum one year onsite comprehensive warranty covering all parts & labours is

essential. This period will start from the date of successful delivery / commissioning
of the instrument.

St The technical bids will be opened individually, and examined for short listing the
vendors, and financial bids of only those vendors who have qualified technically will
be opened and compared.

6. The specifications are clearly mentioned as per Annexure-A and the bidder are
requested to submit Bid only if their offer strictly comply with the given
specifications. Please note that no deviation in the required specifications will be
permitted.

7 The instruments / equipments have to be supplied in standard packaging only at
Delhi Pollution Control Committee (Laboratory) at 4™ Floor, ISBT Bldg., Kashmere
Gate, Delhi-110006.

8. The vendor must deliver the instrument within 30 days from the date of placement of
orders. Failing which EMD amount may be forfeited. If the vendor fails to deliver
any or all the equipment for which tenders are invited, the EMD amount may be
forfeited

9 If the equipments / instruments supplied are found defective / unsatisfactory
condition / short supply / other than specifications in the purchase orders, the same
will have to be replaced at the supplier’s risk & cost. No freight & other charges will
be paid by DPCC.

10.  Any dispute arising out of this purchase shall be under the jurisdiction of the courts
of Delhi.

1. The Competent Authority, DPCC reserves the right to reject any or all the tenders
without assigning any reason.

This issue with the approval of competent authority vide UO No. MS/447 dated

12.09.2011.
@ IR Ju

ADMINISTRATIVE OFFICER
Copy to:-
1. Sr. Scientist (IT):- for uploading on DPCC website please.
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STACK MONITORING KIT

1. General

Stack velocity range 0 to 30 m/sec

Stack Temperature range : 0to 600 °C

Particulate Sampling ! At 6 to 60 Ipm

Filter Paper ( Thimble) : Collection of particulates down to 0.3
micron

Gaseous sampling H At 1to2 lpm callection on a set of
impingers, containing selective reagents

Pitot tube : System should have facilities to

connect.accessories like heated probe
system , cyclone separator etc.

i) Modified S —Type pitot tube shall be fabricated from SS 304 or

equivalent.The construction features should be as CPCB Doc. No. Emission
Regulation (Dec. 1985) Part I11.

i) Calibration certification from repute CSIR or IIT labs should be provided
for each pitot tube.

2. Sampling Probe

Fabricated from SS 304 tube of suitable diameter { not less than 15 mm ID) .

The lengths of the pitot tube the sampling probes shall be decided between the user
and the manufacturer.

3. Nozzles

A set of nozzles fabricated from SS 304 or equivalent material with internal
diameters suitable to cover the full range of stack velocities. The leading edge of

the nozzle should be sharp and tapered. The minimum internal diameter of the
nozzle should nor be less than 6 mm.

4. Thimble Holder

Filter holders fabricated from SS 304 suitable to hold cellulose/glass fibre
thimbles.

5. Thermocouple

Thermocouple sensor with analog digital dial gauge capable of measuring

temperature from 0 to 600 ° C covered with stainless steel or mild steel casing with
acid resistant treatment.
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A pair of male/female flanges fabricated out of mild steel with proper hole

for mounting thermocouple sensor, sampling tube and pitot tube.
=

7. Panel Box Sides

Baked stove- enamel finish . It should have suitable arrangements or
housing stop-watch , manometer » rotameter , dry gas meter etc.

8. Back Panel

Hinged door panel of mild steel to contain cold box with 5 impingers qr
cold box can be proved separately.

9. Inclined-cum-Vertical Manometer

Fabricated out of solid acrylic sheets/ blown glass. Inlet and outlet

provided at the ends for filling in gauge fluid, spirit level attached for leveling.
Velocity range: 0 to 30 m/sec.

10. Rotameter

0 to 60 Ipm for particulate monitoring and 1 to 6 Ipm for gaseous
monitoring.

11. Stop —Watch

0 to 60 minutes, one second readout with hold facility.

12. Impingers

Four numbers of 120 ml and two numbers of 250 ml capacity. Facility
should be there for keeping ice at the bottom of impinger box.

13. Vacuum Pump

Rotary design, with a capacity upto 120 Ipm gas flow with single phase
motor, 230 + 10 V. The pump will also have a moisture trap, air inlet valve and
mounted inside a pump housing and should be portable.

14. Dry Gas Meter

The sampling train shall have a dry gas meter with the facility for
measuring temperature and static pressure . The capacity of the meter should be

adequate to record up to 100 Ipm of airflow and a minimum readout of 0.001 cubic
meters. Dry Gas Meters shall be suited for gaseous sampling rates also.
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Mild Steel case with oven-baked stove —enamel finish and ON/OFF switch
with indicator lights.
.

16. Tools

A kit containing the essential tools required for connecting various
components shall be provided with the equipment.

17. Train Leakages

The sampling train after having set up will be tested for leakages by
plugging the inlet . The rotameter shall not give a reading beyond 5 lpm when the
flow has been set at 60 Ipm, also the dry gas meter should give a reading of less
than 5 percent of the air flow.




